# NIST SRM 2135c EPMA Measurements

This study of the SRM 2135c depth profile standard from NIST that consists of alternating layers of Cr and Ni on a Si substrate. The purpose of the study was to test the quantitative capabilities of the electron microprobe with thin films for quantifying thin film compositions of replicate deposited layers on Si substrates.

This study used two software packages for acquisition and analysis of the measured x-ray intensities, Probe for Windows-Enterprise Edition and StrataGem software packages. The SRM is described in detail here:

http://patapsco.nist.gov/srmcatalog/certificates/2135c.pdf

This analysis was performed by assuming a homogeneous mixture of the elements in the film on top of the Si substrate and iterating on the “average” composition. This assumption, although not physically accurate, seems to give the most accurate composition of the total replicate film composition for many systems over a range of replicate layer thickness’. This particular test, using a test specimen with only a few, very thick replicate layers, is an extreme test of this assumption.

This work was done at the CAMCOR Microanalytical Facility at the University of Oregon. Please contact John Donovan for more information. (541) 346-4632, donovan@uoregon.edu

1.) Thin Film (traverse 1)- homogeneous model (all elements in one layer on Si) (10, 15, 20 25 keV)

Note first that analysis of oxygen gave approximately 0.59 wt. % (~2% atomic) in homogeneous model and second, that the homogeneous physical model assumption starts to fail at accelerating voltages below 15 keV due to the extreme thickness of the film (no Si substrate signal could be measured below this voltage).

Even so, the calculated homogeneous composition agrees well with the certified values as a ratio (see table).  The calculated total was 97.23 wt % for the data including the 10 keV data and 99.73 without the 10 keV data (see next page).

Homogeneous Model

| Total Thickness | Cr | Ni | O | Cr | Ni | O |
| --- | --- | --- | --- | --- | --- | --- |
| Certified | 206.3 ug/cm2 | 197.4 ug/cm2 |  | .5110 | .4889 |  |
| Measured | 217.4 ug/cm2 | 206.7 ug/cm2 | 2.5 ug/cm2 | .5084 | .4857 | .0059 |

Thin Film (traverse 1 again)- homogeneous model (all elements in one layer on Si) (15, 20 25 keV)

Same as above but without 10 keV data.

Homogeneous Model (without 10 keV data)

| Total Thickness | Cr | Ni | O | Cr | Ni | O |
| --- | --- | --- | --- | --- | --- | --- |
| Certified | 206.3 ug/cm2 | 197.4 ug/cm2 |  | .5110 | .4889 |  |
| Measured | 216.4 ug/cm2 | 205.0 ug/cm2 | 2.37 ug/cm2 | .5106 | .4838 | .0056 |

These measured compositional values agree excellently with the “certified” values for Ni and Cr (relative errors within .08% for Cr and within 1.1% for Ni), although the total calculated mass thickness was high by approximately 4.9% for Cr and 3.8% for Ni.

EXPERIMENTAL DATA

Element	Line	Standard	HVstd	HVx	Ix/Istd	Kratio	Iter

O	Ka	STD-912	10.00	10.00	0.0207	0.0055	x

15.00	15.00	0.0177	0.0037	x

20.00	20.00	0.0165	0.0028	x

25.00	25.00	0.0159	0.0023	x

----------------------------------------------------------------------

Ni	Ka	STD-528	10.00	10.00	0.3984	0.3635	x

15.00	15.00	0.4709	0.4611	x

20.00	20.00	0.3848	0.3810	x

25.00	25.00	0.2733	0.2717	x

----------------------------------------------------------------------

Cr	Ka	STD-524	10.00	10.00	0.5636	0.5302	x

15.00	15.00	0.4685	0.4547	x

20.00	20.00	0.3704	0.3624	x

25.00	25.00	0.2657	0.2609	x

----------------------------------------------------------------------

Si	Ka	STD-514	10.00	10.00	0.0008	0.0008	x

15.00	15.00	0.0161	0.0159	x

20.00	20.00	0.0650	0.0645	x

25.00	25.00	0.1122	0.1116	x

======================================================================

USED STANDARDS

STD-912

Tilt angle            :	  0	degrees

Layer #	Element	Concentration		Mass Depth	Thickness	Density

(# atoms)	(µg/cm²)	(nm)

1	C	1.0000	1.0000	 4.520	  20.0	   2.3

----------------------------------------------------------------------

Substrate	Mg	0.6023	0.4996

O	0.3963	0.4997

Ca	0.0014	0.0007

----------------------------------------------------------------------

STD-528

Tilt angle            :	  0	degrees

Layer #	Element	Concentration		Mass Depth	Thickness	Density

(# atoms)	(µg/cm²)	(nm)

1	C	1.0000	1.0000	 4.520	  20.0	   2.3

----------------------------------------------------------------------

Substrate	Ni	1.0000	1.0000

----------------------------------------------------------------------

STD-524

Tilt angle            :	  0	degrees

Layer #	Element	Concentration		Mass Depth	Thickness	Density

(# atoms)	(µg/cm²)	(nm)

1	C	1.0000	1.0000	 4.520	  20.0	   2.3

----------------------------------------------------------------------

Substrate	Cr	0.9882	0.9688

Fe	0.0028	0.0026

O	0.0090	0.0287

----------------------------------------------------------------------

STD-514

Tilt angle            :	  0	degrees

Layer #	Element	Concentration		Mass Depth	Thickness	Density

(# atoms)	(µg/cm²)	(nm)

1	C	1.0000	1.0000	 4.520	  20.0	   2.3

----------------------------------------------------------------------

Substrate	Si	1.0000	1.0000

======================================================================